Warsaw, 27.09.2022 r.

REQUEST FOR PROPOSAL (RFP) NO. Z/355/BS/09/22

Detailed order description

- a. Object of the order

The subject of the order is purchase and delivery of new set of motorized mechanical positioning stages with drivers and assembly base.

The components of the set being the subject of order have to satisfy minimum requirements described below:

| Item | Minimum technical parameters |
| --- | --- |
| 1. Vertical positioning axis | Travel range: 200 mm Width: min. 200 mm Positioning accuracy*: ± 1 m Bidirectional repeatability*: ± 0.5 m Pitch/Yaw: 27/5 arcsec Max. speed: 100 mm/s Max. load: 40 kg (vertical) Dust protection, mechanical break, optical encoder, cable management, dual feedback loop |
| 2. Planar XY stage | Travel range X: min. 120 mm Travel range Y: min. 120 mm Overall dimensions: max. 400 x 400 x 100 mm (W x L x H) Moving platform dimensions: min. 200 x 200 mm (W x H) Positioning accuracy X*: ± 0.5 m Positioning accuracy Y*: ± 0.5 m Bidirectional repeatability X*: ± 0.15 m Bidirectional repeatability Y*: ± 0.15 m Pitch/Yaw X: < 6.2 / 6.2 arcsec Pitch/Yaw Y: < 6.2 / 6.2 arcsec Max. speed X: 0.8 m/s (@5kg external load) Max. speed Y: 1.3 m/s (@5kg external load) Max. acceleration X: 8 m/s2 (@5kg external load) Max. acceleration Y: 13 m/s2 (@5kg external load) Load of XY system: > 10 kg Encoder: optical with 25 nm resolution Orthogonality between X and Y axes: <5 arcsec Cable management |
| 3. Report consisting of measurement results of the axes performance after assembly on a base in target configuration. | Yes |
| 4. Dedicated drivers and controllers for stages. These elements are necessary to control the motion of mechanical positioning elements and allow to obtain the desired positioning accuracy and dynamics of movement. Components assembled in standard 19” rack cabinet. Three additional driver slots free for future system extension. | Yes |
| 5. a) Laser control module providing laser triggering. b) Electronic module for synchronous work of stages and galvoscanners (enabling infinite field of view function). | Yes |
| 6. Power supplies and wiring necessary for the startup and operation of mechanical positioning elements. | Yes |
| 7. Assembly base and components necessary for fixing the positioning axes. Base is critical to obtain the required positioning parameters such as positioning accuracy, dynamic performance with load. Motorized stages mounted on base. | Base material: granite Flatness of mounting surfaces : <6 m / m Orthogonality of vertical axes to XY plane: <5 arcsec Total height of the portal: min. 750 mm The width of granite base min.: 900 mm The length of granite base min.: 900 mm Thickness of granite components > 100 mm Additional holes in granite (beside stage mounts): - general granite plate: 14 pcs (M6) - for vertically mounted bredboards 8 pcs (M6) - for laser mounting 4 pcs (M6) Optical breadboards with M6 hole pattern mounted to vertical granite planes: 2 pcs 200 x 200 x 15 mm (aluminum, anodized) |
| 8. Upper breadboard | Custom shape with beam aperture and M6 hole pattern Thickness min. 50 mm Min. width 900 mm Min. length: 750 mm |
| 9. Granite supports | 4 legs, passive vibration isolation, material: steel |

*- positioning accuracy, repeatability defined for entire travel range

        - The warranty period for the device is minimum 12 months.

        - The Contractor is responsible for the equipment until the device is collected by the Ordering Party

        - The Contractor undertakes to complete installation and commissioning of the device (possible remote installation and commissioning)

        - The Contractor will make final tuning of the system after delivery by means of online/remote support. The Contractor will provide general technical online/remote support for 2 years after delivery.

        - The Ordering Party, prior to selecting the best offer, reserves the right to request the delivery (or for inspection) of original documents confirming that the criteria required by this specification and national and European Community law are met

      - Type of order - delivery

      - CPV code: 38000000-5 Laboratory equipment, optical,  precise;

      - Is it permissible to submit a partial offer: NO

      - Is it permissible to submit a variant offer: NO

      - Place of execution of the order: Fluence Sp. z o.o. - Wrocławski Park Biznesu, ul. Bierutowska 57-59, 51-317 Wrocław

      - Procedure: RFQ

- Ordering Party: Fluence sp. z o.o, ul. Kolejowa 5/7, 01-217 Warsaw, NIP: 5272776154, REGON: 365029156 entered in the Register of Entrepreneurs of the National Court Register maintained by the District Court for the capital city. Warsaw in Warsaw, XII Commercial Division of the National Court Register under KRS number 0000629831

- Other information:

The Order is executed under the project: " High-precision micromachining station based on improved femtosecond laser " under Activity 1.1. "Enterprise R & D projects", Sub-activity 1.1.1: "Industrial research and development works by enterprises", Operational Program Intelligent Development 2014-2020 co-financed by the European Regional Development Fund, POIR Contract No. 01.01.01-00- 1315 / 19-00

- Terms of participation in the proceedings:

  - Contractors involved in the procedure must be in an economic and financial situation ensuring timely and compliant with requirements performance of the order. In addition, they must have the necessary knowledge, experience and technical potential, and have the ability to perform the Order.

  - Incomplete offers and offers, which will be submitted after the expiry date set out in this RFP , will not be considered.

- Criteria and method of evaluation of tenders

  - Criteria for evaluation of tenders

Net price - 100% (max 100 points)

  - How to evaluate offers

Points for the above criterion will be calculated according to the following formula:

P = (𝐶min / Co) 𝑥 [100]

where:

P - the number of points awarded to the Contractor for the offered price

Cmin - the lowest offered price in proceedings

Co - price offered in the evaluated offer.

The maximum number of points to earn is 100. The tender that receives the highest number of points will be the most advantageous.

The Ordering Party informs that in order to standardize currencies, when evaluating offers, the values given in a currency other than PLN will be converted using the average exchange rate announced by the National Bank of Poland on the date of the submission of offer.

6. Preparation and delivery of the offer

- The complete, signed offer should be submitted on the Offer Form (under pain of exclusion from the procedure) constituting Appendix 2 to this Inquiry by e-mail to the following address: biuro@fluence.pl or in paper form to Fluence sp. z o.o., Kolejowa 5/7, 01- 217 Warsaw by 05.10.2022.

- Contact person for technical questions: Bogusz Stępak, e-mail: bstepak@fluence.pl

- The submitted offer is binding to the Contractor for 60 days from the deadline for submission of tenders.

- In the title of the email or on the envelope with the offer, please enter the order number i.e. Z/355/BS/11/20

- If the prices on the offer are expressed in a currency other than Polish zloty, they will be converted using the average selling rate announced by the National Bank of Poland, applicable on the date of the offer,

- The cost of preparation and submission of the offer is borne by the Contractor.

- List of required statements and documents confirming the fulfillment of participation in the procedure by Contractors and other documents:

The Ordering Party requires to present the following documents:

  - Offer form - attached as Appendix 2 to this inquiry;

  - Declaration that the bidder has no relations with the Ordering Party - the template of the declaration is attached as Appendix 3 to this inquiry;

  - A copy of the power of attorney for the person / persons signing / signing the offer to incur obligations on behalf of the Contractor submitting the offer, if the authorization is not disclosed in the National Court Register or CEiDG (not applicable to foreign entities);

  - The offer containing the catalog part numbers with prices.

- Schedule of the contract

The order fulfillment date (single delivery) is provided by the Contractor in the offer Form, however, cannot be later than 12 weeks from the date of signing the contract.

- Additional conditions

- The offer will remain unchanged throughout the duration of the contract.

- In the course of examination and evaluation of the offers, the Ordering Party may demand explanations regarding the content of the submitted offers from the tenderers.

- The Ordering Party shall award the contract to the Contractor whose offer meets all the requirements set out in the inquiry and presented the most advantageous offer based on the selection criterion specified in this inquiry

- The Ordering Party does not provide for reimbursement of the costs of participation in the procedure.

- Offers submitted after the deadline will not be considered.

- The Ordering Party reserves the right to change the terms and conditions set out in this RFP, which will be communicated via the appropriate websites and portals. In the event of such changes, the Purchaser may change the date for submission of tenders.

- The Ordering Party also reserves the right to revoke or close the proceedings at any stage without giving a reason.

- The Ordering Party also reserves the right to make significant changes to the provisions of the contract in relation to the content of the offer on the basis of which the Contractor was selected, provided that these changes are dictated by circumstances that may affect the proper implementation of this order. A written form is required to amend the contract.

- The Ordering Party allows for the rejection of tenders whose substantive scope or grossly low price will raise reasonable doubts.